## Sheet1
소프트 밸리데이션
3 | 소프트웨어 도식
| 3-1 | Block diagram
| 3-2 | Program Language
|  |  | SN8P2707AQ | C LANGUAGE
|  |  | SN8P2602BS | ASSEMBLER
| 3-3 | Software Platform(설계도)
| * | SN8P2707AQ
|  | 변수 | 자료명 | 범위 | 형식
|  | FullSign | data | 0x33~0x43 | 문자
AutoSign
LevelLowUsed
Tc0rBuf
adc_result
debug_check
MaxState
MuteSign
BeepCnt
Power2602
LockSign
CurrentState
AllState
two_beep_flag
aux_counter
aux_flag
music_n
|  | Flag2p5ms |  | 0x4F
|  | KeyPER1 |  | 0x51~0x53
KeyValue2
cartridge_count
|  | cartridge_hcount |  | 0x0C
|  | cartr_hcount1 |  | 0x0E~0x13
cartr_hcount2
cartr_hcount3
level_count
beep_set
debug_count
|  | Second0p5Flag |  | 0x1D~0x25
Second0p5Flag2
Second0p5Counter
Minute
KeyValue
FlashLifeType
Chip2602Stable
ErrorFlag
cartridge_buffer
| * | SN8P2602BS
|  | CurrentKey | data | 0x00~0x24 | ASSEMBLER
LastKey
KeyCount
Second
TimeCount
dataOut
DelayCount
counter
dataIn
UseCounterHigh
UseCounterLow
UseCounterHigh1
UseCounterLow1
FlashTrigTime
KeyValue
Flags
Flags2
Flags3
delay_err
debug_count
skin_delayCounter
level_h_counter
level_counter
life_plus
cartr_counter
auto_delayCount
type_lamp
type_lamp2
type_lamp3
eep_count
last_life_plus
UseCounterHigh2
UseCounterLow2
UseCounterflag
Chip2602Stable
ErrorFlag
cartridge_buffer
|  | FlashKeySign | bit | 0x0E | 바이너리
FullChargeSign
SmallFanSign
TouchSkinSign
Glass1Sign
Glass2Sign
AutoFlashSign
|  | SEND_MAK |  | 0x0F
CHECK_SAK
SIGN_5MS
SIGN_EEP_ADD
|  | f_lamp_max |  | 0x10
debug_flag
auto_1time
f_lamp_check
cartrige_s1
cartrige_s2
f_manual_once
f_erro
|  | SmallFanSign1 |  | 0x11
EPP_MAX
|  | UseCounterR1 equ |  | 0x24
| * | SN8P2707AQ
|  | 변수 | 자료명 | 수용범위 | 한계 | 기본값
|  | FullSign | data | 0x33~0x43 | 000H-17FH | 0x00
AutoSign
LevelLowUsed
Tc0rBuf
adc_result
debug_check
MaxState
MuteSign
BeepCnt
Power2602
LockSign
CurrentState
AllState
two_beep_flag
aux_counter
aux_flag
music_n
|  | Flag2p5ms |  | 0x4F
|  | KeyPER1 |  | 0x51~0x53
KeyValue2
cartridge_count
|  | cartridge_hcount |  | 0x0C
|  | cartr_hcount1 |  | 0x0E~0x13
cartr_hcount2
cartr_hcount3
level_count
beep_set
debug_count
|  | Second0p5Flag |  | 0x1D~0x25
Second0p5Flag2
Second0p5Counter
Minute
KeyValue
FlashLifeType
Chip2602Stable
ErrorFlag
cartridge_buffer
| * | SN8P2602BS
|  | 변수 | 자료명 | 수용범위 | 한계 | 기본값
|  | CurrentKey | data | 0x00~0x24 | 000H-OFFH | 0x00
LastKey
KeyCount
Second
TimeCount
dataOut
DelayCount
counter
dataIn
UseCounterHigh
UseCounterLow
UseCounterHigh1
UseCounterLow1
FlashTrigTime
KeyValue
Flags
Flags2
Flags3
delay_err
debug_count
skin_delayCounter
level_h_counter
level_counter
life_plus
cartr_counter
auto_delayCount
type_lamp
type_lamp2
type_lamp3
eep_count
last_life_plus
UseCounterHigh2
UseCounterLow2
UseCounterflag
Chip2602Stable
ErrorFlag
cartridge_buffer
|  | FlashKeySign | bit | 0x0E | 000H-OFFH | 0x00
FullChargeSign
SmallFanSign
TouchSkinSign
Glass1Sign
Glass2Sign
AutoFlashSign
|  | SEND_MAK |  | 0x0F
CHECK_SAK
SIGN_5MS
SIGN_EEP_ADD
|  | f_lamp_max |  | 0x10
debug_flag
auto_1time
f_lamp_check
cartrige_s1
cartrige_s2
f_manual_once
f_erro
|  | SmallFanSign1 |  | 0x11
EPP_MAX
|  | UseCounterR1 equ |  | 0x24
4 | Software 성능 요구사항
|  | Function |  | Requirement | Design | Control
|  | Key | Operation Mode | Control operation mode | IPL Mode toggle | CPU1,2
|  |  | IPL Pulse width mode more/less | Pulse on time control | 1~125msec(step 1msec) | CPU1,2
|  |  | IPL Pulse delay more/less | Pulse off time control | 5~500msec(step 1msec) | CPU1,2
|  |  | IPL Output power more/less | Output power control | 5~40msec(step 1msec) | CPU1,2
|  |  | IPL Temperature more/less | Temperature control | -5ºC~25ºC (step 1ºC) | CPU1,2
|  | Display | IPL Pulse type | NO Display | Single | CPU1,2
|  |  | IPL Pulse width | NO Display | Range : 1~125 msec | CPU1,2
|  |  | IPL Delay time | NO Display | Range : 5~500 msec | CPU1,2
|  |  | IPL Output power | LED Level indicator | Range : 15~25J / cm2 | CPU1,2
|  |  | operation | Flash lamp on/off | Detect hand/foot switch | CPU1,2
5 | Software 관리기준
| 5-1 | 사용자와 기기간의 중요 관리항목
User Menual inside interface specification ( 푸로맥스 사용설명서 사용 )
| 5-2 | 사용시 기능 요구사항
User Menual inside interface specification ( 푸로맥스 사용설명서 사용 )
| 5-3 | 적용 연산방식
Applied Digital and analog
| 5-4 | 빛의 파장 설계한도 ( 푸로맥스 석영 적용 참고사항 )
|  |  | 미백효과 | 430nm~1200nm ?
|  |  | 여드럼 완화 | 430nm~1200nm ?
|  |  | 기미 완화 | 430nm~1200nm ?
|  |  | 제모효과 | 530nm~800nm ?
| 5-5 | 설계상 기기의 사용한계 (푸로맥스 램프수명 적용 )
Device limitation due td software
| 5-6 | 설계특성의 복구
Recovery characterist
| 5-7 | 내부회로 체크사항
Internal C, or assembler
| 5-7 | 동작시간및 기억장치
|  | Operate time | Normal mode Flash time | Level 1 | 3.5s more/less
|  |  |  | Level 2 | 3.5s more/less
|  |  |  | Level 3 | 3.5s more/less
|  |  |  | Level 4 | 4.5s more/less
|  |  |  | Level 5 | 4.5s more/less
|  |  |  | Level 6 | 4.5s more/less
|  |  |  | Level 7 | 5.5s more/less
|  |  |  | Level 8 | 5.5s more/less
|  |  | Auto Mode Flash time | Level 1 | 5s more/less
|  |  |  | Level 2 | 5s more/less
|  |  |  | Level 3 | 5s more/less
|  |  |  | Level 4 | 5s more/less
|  |  |  | Level 5 | 5s more/less
|  |  | No use auto off time | All | 10 minute
|  | Memory | Device | Memory(RAM) | Memory(ROM)
|  |  | SN8P2707AQ MCU | 256*8bit | 4k*16bit
|  |  | SN8P2602BS MCU | 48*8bit | 1k*16bit
|  |  | 11LC010 EEP ROM |  | 1k bit
9 | 소프트외어 시험(검증)
Unit test & all system test
4-1 | 소프트 인터페이스의 정의
|  | 출력모쥴 입력모쥴 | UP/DOWN BUTTON | 전원 ON/OFF BUTTON | 피부 접촉 감지 센서 | 카드리지 유무센서
|  | LED 기능표시 | 1~8 단계 Level 표시 | Power | x | Error 표시
|  | LED 레벨표시 | 1~8 단계 | x | x | x
|  | LED 동작표시 | x | x | 접촉시 flash standby | x
4-2 | 하드웨어 링크
SN8P2707AQ
SN8P2602B
EEP ROM_11LCO10
|  | 구분 | 링크 |  | 역할
|  | ALU | Program count, Status and Control, Interrupt, ADC, !/O EEPROM, Analog Comparator |  | MCU의 Brain에 해당하며 값들의 계산과 논리비교 데이터 변환의 역할을 담당
|  | TIMER | Program count |  | 정해진 영역에서 카운터를 증가하거나 감소
|  | I/O | 1/o Module 1~7 |  | MCU의 외부에서 들어오는 신호를 감지하고 입력받음
|  | A/D | Analog Comparator |  | 외부의 아나로그 신호를 디지털 신호로 변환
|  | EEP ROM | ALU&EEP ROM |  | 외부에서 입력된 값을 저장
|  | Flash Memory | ALU&EEP ROM |  | 펌웨어의 Hex 코드를 저장
통합 시스템 시험
| * | Set-up
|  | 테스트 케이스 ID | MAIN SET & PROBE SET ( 타사 참고 자료 SET-0001 )
|  | 목적 | MCU의 포트, 타이머, ADC, 초기화 함수와 변수 데이터, 초기화 프로그램 메모리 명령에 따라 변수를 불러들여 초기화 한다
|  | 테스트 조건 | 대상 변수들의 값이 초기값과 다른 값을 가지고 있다
|  | 단계 | 확인대상 |  | 예상출력값 | 실행결과
|  | 1 | SN8P2602 | P0 | 0x01 | 0x01
|  | 2 |  | P0M | 0x01 | 0x01
|  | 3 |  | P0UR | 0x00 | 0x00
|  | 4 |  | P1 | 0x00 | 0x00
|  | 5 |  | P1M | 0x06 | 0x06
|  | 6 |  | P1UR | 0x01 | 0x01
|  | 7 |  | P1W | 0x00 | 0x00
|  | 8 |  | P5 | 0x01 | 0x01
|  | 9 |  | P5M | 0x03 | 0x03
|  | 10 |  | P5UR | 0xFC | 0xFC
|  | 11 |  | P5DR | 0x00 | 0x00
|  | 12 |  | OSCM | 0x00 | 0x00
|  | 13 |  | T0M | 0xA0 | 0xA0
|  | 14 |  | T0C | 0x64 | 0x64
|  | 15 |  | INTEN | 0x10 | 0x10
|  | 16 |  | INTRQ | 0x00 | 0x00
|  | 17 |  | STKP | 0x83 | 0x83
|  | 18 |  | WDTR | 0x5A | 0x5A
|  | 19 |  | PCL | 0x00 | 0x00
|  | 20 |  | PCH | 0x00 | 0x00
|  | 21 |  | STK0L | 0x00 | 0x00
|  | 22 |  | STK0H | 0x00 | 0x00
|  | 23 |  | STK1L | 0x00 | 0x00
|  | 24 |  | STK1H | 0x00 | 0x00
|  | 25 |  | STK2L | 0x00 | 0x00
|  | 26 |  | STK2H | 0x00 | 0x00
|  | 27 |  | STK3L | 0x00 | 0x00
|  | 28 |  | STK3H | 0x00 | 0x00
|  | 29 |  | BZM | 0x00 | 0x00
|  | 30 |  | TC0C | 0x00 | 0x00
|  | 31 |  | TC0M | 0x00 | 0x00
|  | 32 |  | TC0R | 0x00 | 0x00
|  | 33 |  | PEDGE | 0x10 | 0x10
|  | 1 | SN8P2707 | P0 | 0xFF | 0xFF
|  | 2 |  | P0M | 0xFF | 0xFF
|  | 3 |  | P0UR | 0x00 | 0x00
|  | 4 |  | P1 | 0xFE | 0xFE
|  | 5 |  | P1M | 0xFE | 0x06
|  | 6 |  | P1UR | 0x01 | 0x01
|  | 7 |  | P1W | 0x00 | 0x00
|  | 8 |  | P2 | 0xF8 | 0xF8
|  | 9 |  | P2M | 0xF8 | 0xF8
|  | 10 |  | P2UR | 0x07 | 0x07
|  | 11 |  | P4 | 0xE0 | 0xE0
|  | 12 |  | P4M | 0xE4 | 0xE4
|  | 13 |  | P4UR | 0x19 | 0x19
|  | 14 |  | P4CON | 0x02 | 0x02
|  | 15 |  | P5 | 0x15 | 0x15
|  | 16 |  | P5M | 0xBF | 0xBF
|  | 17 |  | P5UR | 0x40 | 0x40
|  | 18 |  | OSCM | 0x00 | 0x00
|  | 19 |  | T0M | 0xC0 | 0xC0
|  | 20 |  | T0C | 0x64 | 0x64
|  | 21 |  | INTEN | 0x10 | 0x10
|  | 22 |  | INTRQ | 0x00 | 0x00
|  | 23 |  | STKP | 0x87 | 0x87
|  | 24 |  | WDTR | 0x00 | 0x00
|  | 25 |  | PCL | 0x00 | 0x00
|  | 26 |  | PCH | 0x00 | 0x00
|  | 27 |  | STK2L | 0x00 | 0x00
|  | 28 |  | STK2H | 0x00 | 0x00
|  | 29 |  | STK3L | 0x00 | 0x00
|  | 30 |  | STK3H | 0x00 | 0x00
|  | 31 |  | STK4L | 0x00 | 0x00
|  | 32 |  | STK4H | 0x00 | 0x00
|  | 33 |  | STK5L | 0x00 | 0x00
|  | 34 |  | STK5H | 0x00 | 0x00
|  | 35 |  | STK6L | 0x00 | 0x00
|  | 36 |  | STK6H | 0x00 | 0x00
|  | 37 |  | STK7L | 0x00 | 0x00
|  | 38 |  | STK7H | 0x00 | 0x00
|  | 39 |  | DAM | 0x00 | 0x00
|  | 40 |  | TC0M | 0xD1 | 0xD1
|  | 41 |  | TC0C | 0x00 | 0x00
|  | 42 |  | TC0R | 0xFF | 0xFF
|  | 43 |  | PEDGE | 0x10 | 0x10
|  | 44 |  | TC1M | 0x60 | 0x60
|  | 45 |  | TC1C | 0x00 | 0x00
|  | 46 |  | TC1R | 0x00 | 0x00
|  | 47 |  | ADM | 0xD1 | 0xD1
|  | 48 |  | ADR | 0x80 | 0x80
|  | 49 |  | P1OC | 0x00 | 0x00
|  | 50 |  | SIOM | 0x00 | 0x00
|  | 51 |  | SIOR | 0x00 | 0x00
User Input
|  | 테스트 케이스 ID |  | UI-0001
|  | 목적 | 기능 스위치의 신호를 감지한다
|  | 테스트 조건 | 기능 스위치 신호의 Bit를 변화시킨다
|  | 단계 | 확인대상 |  | 예상출력값 | 실행결과
|  | 1 | SN8P2707 | 0xF7 | 0xF7 | 0xF7
|  | 2 |  | 0xFD | 0xFD | 0xFD
|  | 3 |  | 0xFA | 0xFA | 0xFA
|  | 4 |  | 0xF7 | 0xF7 | 0xF7
|  | 5 | SN8P2602 | 0xEF | 0xEF | 0xEF
User Input
|  | 테스트 케이스 ID |  | UI-0002
|  | 목적 | 레벨 스위치의 신호를 감지한다
|  | 테스트 조건 | 레벨 스위치 신호의 Bit를 감지한다
|  | 단계 | 확인대상 |  | 예상출력값 | 실행결과
|  | 1 | SN8P2707 | 0xF7 | 0xF7 | 0xF7
|  | 2 |  | 0xFE | 0xFE | 0xFE
|  | 3 |  | 0xFB | 0xFB | 0xFB
User View
|  | 테스트 케이스 ID |  | UV-0001
|  | 목적 | LED 표시에 데이터를 전송한다
|  | 테스트 조건 | 램프충전 전압을 변화시키며 입력한다
|  | 단계 | 확인대상 |  | 예상출력값 | 실행결과
|  | 1 | mode | FLASH | FLASH | FLASH
|  |  | SN8P2602 Data | 0x20 | 0x20 | 0x20
User View
|  | 테스트 케이스 ID |  | UV-0002
|  | 목적 | 레벌을 선택하여 DISPLAY 한다
|  | 테스트 조건 | LEVEL 값을 변화시키며 입력한다
|  | 단계 | 확인대상 |  | 예상출력값 | 실행결과
|  | 1 | SN8P2707 | 0xFE | 0xFE | 0xFE
|  | 2 |  | 0xFC | 0xFC | 0xFC
|  | 3 |  | 0xF8 | 0xF8 | 0xF8
|  | 4 |  | 0xF0 | 0xF0 | 0xF0
|  | 5 |  | 0xE0 | 0xE0 | 0xE0
|  | 6 |  | 0xC0 | 0xC0 | 0xC0
|  | 7 |  | 0x80 | 0x80 | 0x80
|  | 8 |  | 0x00 | 0x00 | 0x00
7-3-1 |  | Software Architecture Design Chart (SAD)
| 7-1 | Software Diagram (Device Firmware)
| 7-1-1 | SN8P2707 FLOW CHART
7-1-1-A
7-1-1-B
| 7-1-2 | SN8P2602B FLOW CHART
| 7-2 | Software diagram description
The software consists with main processor, LED signal data controller and data storage.
7.2.1 Main processor
1) Level Battery Charge Check: Check the Level of battery charge (3 Level).
2) Key Input Control: Check the key input pulse and determine the key function
3) Power On/Off: System power on/off and the initial setting
4) Mode Selection: Select 25beat measuring mode or 3M measuring mode
5) Start/Stop operation: Read the key input control signal and LED signal Data controller signal or Read the Communication with PC signal and LED signal Data controller signal
6) Data Calculation: Calculation of Heart rate and Heart pulse
7) Display module: Display Battery level, selected mode, calculated data (Heart Rate, Heart Pulse) and number of stored records
8) Delete the data: Delete the stored data
9) Communication with PC: Communication with PC and the system
7.2.2 LED signal data Controller
1) Check the finger in/out and converter the LED signal to data
7.2.3 Data storage management
1) Store or delete the calculated data (Heart rate and Heart pulse)
| 7-3 | Software Architecture Design Specification (Device Firmware)
7.3.1 Initialization
|  | Identification | Function | Description
|  | Initialization key input | Power On | 1. power&mute switch input
2. Read Key input pulse time (Long pulse) check
* Power standby
3. The Same time +&- Level switch input
4. Read key input same time check and (long pulse) input
|  |  | Power Off | 5. power&mute switch input
6. Read key input pulse time ( long pulse ) check
|  |  | Mute | 8. power&mute switch input
9. Read key input pulse time (Short time) check
|  |  | Auto shot | 1. Auto switch input
2. Read key input pulse ( Normal time ) check
|  |  | Level selection | 1. Level +,- switch input
2. Reade Key Input pulse time (Short pulse) check
|  | Identification | Function | Description
|  | Initialization structure | cartridge err check | 1. IC11(SN8P2602BS) PIN8 and pin 9 HIGH LEVEL CHECK
|  |  | Flash lamp Life time Memory | 1. IC16(11LC010) EEP ROM Memory check
|  |  | timer seting | 1. key Use time check
2. capacitor charge time check
3. Flash after charge time check
|  |  | skin touch | 1. skin sensor SN8P2602 PIN10 input Voltage check
|  |  | Fan err | 1. IC8(SN8P2707) PIN22 LEVEL CHECK
2. IC11(SN8P2602BS) PIN11 LEVEL CHECK
|  |  | Level selection | IC8(SN8P2707) PIN26 VOLTAGE CHECK
|  |  | Auto shot selection | 1, Level 5 Max
2. Skin sensor on check
|  |  | Mode display | 1. LED Display
|  | Identification | Function | Description
|  | Children Safety Protection | Power | Power start After standby and two key the same time press After power ON
|  |  | Flash | No touch Skim no flash.
|  |  | Cartridge check | Cartridge not insert no flash
|  | Identification | Function | Description
|  | Energe Level Control | Normal Level | IC8(SN8P2707) Pin25 voltage level check control (Level 1~8)
|  |  | Auto Level | IC8(SN8P2707) Pin25 voltage level check control (Level 1~5)
|  |  | Cartridge check | IC11(SN8P2602) Pin8,9 voltage level check (1~5, Level 1~3)
|  | Identification | Function | Description
|  | Auto Function | Level Check | IC8(SN8P2707 pin25) voltage level (Level 1~5)
|  |  | Cartridge check | IC11(SN8P2602 pin8,9) voltage level (1~5, Level 1~3 Level )
|  | Identification | Function | Description
|  | Shot Function | chagge voltage check | IC8(SN8P2707) pin25 voltage Level Full charge check (Level 1~8)
|  |  | skin touch check | IC11(SN8P2602) PIN10 voltage level check
|  | Identification | Function | Description
|  | Mute Function | Power / Mute switch | IC8(SN8P2707) pin34 off
|  | Identification | Function | Description
|  |  | Memory count check | Memory check (IC16 11LC010 EEP ROM)
|  | Shot flash count Function | Count memory | Charge Level check (Level 1~3) One time Use memory
Charge Level check (Level 4~5) Three time Use memory
Charge Level check (Level 6~8) five time Use memory
|  |  | Max use time memory check | Max shot Memory check (Max 30, 000 time use)
|  | Identification | Function | Description
|  | Auto power off Function | Auto charge off | Fan err
Cartridge err
Charge over time err
Shot memory over err
|  |  | Auto power off | Key use of the Lapse time ( All key ten minute ) err
| 8. | Software Design Specification (SDS)